## Slide 1
Mirko Belliato, MD
Chief of the Advanced Respiratory Intensive Care Unit
U.O.C. Anestesia e Rianimazione 1
Fondazione IRCCS Policlinico San Matteo, Pavia, Italy
Member of EuroELSO Steering Commitee
New Concepts in ECLS

## Slide 2: Disclosures last 5 yrs:
Advisor, consultant and congress speaker for Hamilton Medical AG, CH
Advisor, consultant and congress speaker for Eurosets srl, Italy

## Slide 3: Summary
New monitoring on ECMO performance, patient’s organs recovery and patient-ECMO interaction
Hybrid ECMO
New ECMO devices (multi-purpose, from OR to transport)

## Slide 4: We can describe the patient metabolic condition as:
Oxygen need
CO2 production

## Slide 5: In a patient with stable metabolic conditions: The total intake must be equal to the O2 needs
Oxygen delivery = V’O2 [ml/min ]
VO2ML
VO2NL
Consumption of O2
Notes: We know that in stable metabolic conditions the oxygenation power is the sum of the VO2 ML and VO2 natural that is an indicator of the consumption of the oxygen by the patient

## Slide 6
Landing® Real Time Monitor, EUROSETS S.r.l. - Medolla (MO), Italy
How to measure continuously the V’O2 ML ?

## Slide 7
V’O2ML on line monitoring
V’O2ML monitored is based on oxygen content approximation obtained from continuous measurements of saturation and haemoglobin with not invasive infrared/spettrophotometry method.

## Slide 8: Oxygenator aging: V’O2ML
Oxygen transfer from membrane lung (V’O2ML) during the entire period of use of the first oxygenator, an ECMO Adult Oxygenator (by Eurosetstm). Actual values are plotted together with the factory theoretical relationship between V’O2ML and blood flow
Belliato et al. J Clin Monit Comput DOI 10.1007/s10877-016-9927-4
Blood Flow [L/min]

## Slide 9
CO2 removal = V’CO2 [ml/min]
VCO2ML
VCO2NL
In a patient with stable metabolic conditions:
The total CO2 removal must be equal to the CO2 production
CO2 production
Notes: At a stable metabolic condition the total Co2 removal is the sum of the VCO2 ML and VCO2 natural lung that is an indicator of the production of CO2 of the patient

## Slide 10: V’CO2ML measurement
V’CO2ML = (ctCO2IN - ctCO2OUT) x 10 x BF
where:

## Slide 11: V’CO2ML not invasive measurement

## Slide 12
Landing® Real Time Monitor, EUROSETS S.r.l. - Medolla (MO), Italy
How to measure continuously the V’CO2 ML?

## Slide 13
“..The volumetric capnometer used in conjunction with the ECMO oxygenator is reliable for continuous monitoring of V′CO2 and useful for oxygenator dead space calculation..”

## Slide 14
Index | Units | Ranges | Meaning
GF | l/min | 0,0-20,0 | Measured sweep gases
eCO2% | % | 0,0-8,0 | CO2 concentration in exhausted gases
Landing® Real Time Monitor, EUROSETS S.r.l. - Medolla (MO), Italy

## Slide 15
Chart: 
| 3 | 3.5 | 3.75 | 3.8 | 3.5 | 3.5 | 2.5 | 3.5 | 3 | 3.5 | 3.5 | 3.5 | 3 | 3 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | 3.5 | 3.5
Oxy Performance | 177 | 175 | 195 | 186 | 150 | 192 | 150 | 183 | 168 | 158 | 151 | 158 | 150 | 150 | 120 | 137 | 133 | 115 | 115 | 103 | 157 | 152
ECMO Blood flow
V’CO2ML vs B.F.ECMO
V’CO2ML (ml/min)
V’CO2 as Index for Aging
Blu curve:
Theorical V’CO2ML with
BF : GF =1:1
of Eurosets ECMO oxy

## Slide 16: Partitioning of V’CO2 during ECMO:

## Slide 18: A Full Support Systems: the hybrid ECMO
The simultaneous use of the two ECMO configurations may represent a novel strategy of advanced treatment to be applied on selected groups of patients such as: ARDS patients supported by VV ECMO suffering from concurrent cardiac failure, typically right-heart dysfunction.

## Slide 19: When do we use the V-VA ECMO?
ARDS severe with bi-ventricular cardiac failure
ARDS severe with left ventricular cardiac failure
ARDS severe with septic shock
Bi-ventricular failure

## Slide 20: Hybrid ECMO configurations

## Slide 22
Flow matching:

## Slide 23: How to describe your ECMO?

## Slide 24
https://ecls.eurosets.com/

## Slide 25: How to select an appropriate ECMO device?
SOURCE: ELSO guideline for ECMO transport, last update May 2015
Best technology available (pumps, disposable, monitor, etc)
Device safety in case system fails (e.g. backup and battery)
Oxygenator adequate performance
Circuit flexibility
User-friendliness, easy device management, easy interface
Built in blood sensors and transducers
Less resources (intra hospital efficiency optimization)
Notes: Semlificare il visual

## Slide 26: Centrifugal pumps available till now:
RotaFlow
Medos DP3
Medtronic BIO-Pump: Affinity, BPX-80 and BP-50
Sorin Revolution 5
Terumo CAPIOX SL

## Slide 27: Magnetic leviation
Belliato 2019

## Slide 28: Advantages using Magnetic LEVITATION Centrifugal Pump
  Magnetically levitated centrifugal pumps require no direct contact between the impeller and the pump housing, with following advantages:
      eliminating particle or heat generation,
      reducing hemolysis,
      reducing thrombus formation,
      lowering risk of mechanical failure (no bearing).
SOURCE: ECMO in Adult Patient, A.Vuylsteke et al. Cambridge uni press 2019
Belliato 2019

## Slide 29
Belliato 2019

## Slide 30: Magnetic LEVITATION Centrifugal Pump
Belliato 2019
No bearings

## Slide 31
Lab data with bovine blood
ECMOlife
Rotaflow
FreeHb in vitro test: ECMOlife vs Rotaflow pump
Belliato 2019

## Slide 32: FLUID DYNAMIC OF ECMOlife
FLUID VELOCITY
PRESSURE DISTRIBUTION
The result confirms the absence of areas with excessive speed , hence:
Low shear stress generation
Low Hemolysis
Low heat generation
ABSENCE OF STAGNATION AREAS
Souce: CFD test
Belliato 2019

## Slide 33: Could we merge the “best” technology with in portable system?

## Slide 34: Intra-Hospital transport
CT SCAN
CATH Lab
Belliato 2019

## Slide 35: INTER-HOSPITAL TRANSPORT

## Slide 36: Circuit Fixing Plate
Aluminum plate
Flexible: possibility to fix the plate on:
    Back side
    Right side
    Left side
    Trolley (any sides)
    Any bar 25x10mm
1
2
3
Fixing lever
Fixing pins
Aluminum plate
Single use organizer
Knob to fix the pump driver
1
2
3
Adult ECMO
Belliato 2019

## Slide 37: HOT POINTS OF A NEW ECMO SYSTEM
Fully magnetic levitation technology
Back up unit integrated
Embedded sensors: Pdrain, Pin, Pout, SvO2, Hb, Tven
Dedicated ECMO set
14 days validated , PC coated
Dedicated trolley for inter-Hospital use fitting to any stretchers.
External battery power bank
Belliato 2019

## Slide 38: THANKS !!!!
Belliato 2019